Model QF4032 Spectrum Analyzer

# Text boxes

9KHz ~ 2200MHz

0 ~ 2210MHz

Center-span, start-span.

±(100Hz+freq.span×2%+display freq.×reference Freq. Accuracy) for Freq.≥10KHz after auto-calibration.

20Hz

1KHz~2000MHz (set in 1,2,5 Sequency)

1KHz~2200Mhz (10 to 98) 2digit

0KHz set (zero sweep mod.)

≤±2%

30Hz~1MHz set in 1.3 sequence, setting can be

Made manually or automatically according to the freq. Span setting.

≤±20%

≤20Hz p-p/0.1s    (frequency span≤500KHz)

Frequency:

Range:

Setting range:

Setting mode:

center frequency:

accuracy:

Resolution:

Frequency span (sp):

Setting range:

Accuracy:

Resolution bandwidth (RB):

3dB bandwidth:

Accuracy:

Residual FM:

≤±300Hz/min (Freq. span≤500KHz    after 1-hr warm-up)

≤-80dBc (R.B.W:100Hz and video band width: 1Hz,20KHz away From the signal

10MHz

aging rate ≤±2×10-8/day

temperature characteristics≤5×10-8 (relative ref. osc at 25℃)

10MHz         2 ~ 5Vp-p

display the frequency at the tunable marker

display the difference in frequency between the reference marker

display the received signal frequency at the marker

1Hz,   10Hz,   100Hz

display freq.×REF. Freq. accuracy±(2counts or 20Hz)

-130dBm ~ +20dBm

8 div are displayed on vertical axis when scale is set  to 10dB / div For all other scale. Setting 10 div are displayed on vertical axis

(topline on the CRT is the reference level)

10dB/div  -80 ~ 0dB            5dB/div   -50 ~ 0dB

2dB/div   -20 ~ 0dB            1dB/div   -10 ~ 0dB    relative reference level

10% reference level/div (calibrated in voltage,   unit V)

(After automatic calibration)

≤±0.2dB  (-10dB ~ 0dB R . B . W 100Hz ~ 1MHz)

≤±0.3dB  (-20dB ~ 0dB R . B . W 100Hz ~ 1MHz)

≤±0.5dB  (-50dB ~ 0dB R . B . W 100Hz ~ 1MHz)

≤±1dB    (-70dB ~ 0dB R . B . W 100Hz ~ 100KHz)

≤±0.5dB (100KHz ~ 2GHz) input attenuation is 20 dB and temperature range is 20℃ to 30℃

-100 ~ +20dBm    (for setting resolution of 0.1dB)

70.8μV ~ 2240mV

≤±0.3dB (0 ~ -50dBm)

≤±0.75dB (-70 ~ 20dBm) after automatic calibration at 50MHz and Freq

span is ≤2MHz   (when RBW , VB , ST , ATT are set to AUTO

≤±0.4dB after automatic calibration

≤±0.4dB after automatic calibration

≤-120dBm (R . B . W:300Hz   VB:1Hz    ATT:0dB    f:1MHz ~ 2GHz)

80Db      ( ATT:0dB   f:5 ~ 800MHz)

Stability:

Drift:

Side band noise:

Reference oscillator:

Frequency:

Stability:

EXT. Ref. input:

Marker:

Normal:

Delta:

Count:

Resolution:

Accuracy:

Amplitude:

Measuring range:

## C.R.T display range accuracy

Divisions:

log :

lin:

Scale linerity error:

log:

Frequency response:

reference level accuracy:

setting range:

log:

lin:

Accuracy:

Resolution bandwidth

Switching deviation

Dynamic range:

average noise level:

2nd/3rd order harmonics distortion:

≤-95dBm (Freq>500KHz   ATT:0dB)  ≤-100dBm typically

Display the level at settable marker

Display the difference in level between the settable marker and the ref. marker

1Hz, 10Hz, 100Hz, 1KHz, 10KHz, 100KHz OFF

Can be set manually or automatically select

dBm, dBuv, dBmV, V, dBuv(emf), dBuv/m

200Hz, 9KHz, 120KHz

45ms (6dB BW 200Hz,         f:10 ~ 150KHz)

1ms  (6dB BW 9KHz 120KHz  f:150KHz ~ 1GHz)

500ms(6dB BW 200Hz,        f:10 ~ 150KHz)

160ms(6dB BW 9KHz,         f:150KHz ~ 30MHz)

550ms(6dB BW 120KHz,       f:30MHz ~ 1GHz)

160ms(6dB BW 200Hz   9KHz)

100ms(6dB BW               f:120KHz)

50Ω       VSWR≤1.5 (ATT≥10dB  f≥30KHz)

+25dBm   (ATT≥10dB)±25Vdc

0 ~ 50dB   10dB steps   (manually / auto)

≤±1dB   (100KHz ~ 1.5GHz)

≤±2dB   (1.5 ~ 2GHz)

Manually or automatically 50ms ~ 100s (in 1, 1.5, 2, 3, 5, 7, sequence, range can be selected

According to frequency span. RBW and video BW)

Free-run  Lin,  VIDEO,  SINGLE,  EXT  TRIGGER

Normal: entire range sweeping

Zone sweep; only zone marker is swept

5-inch electromagnetic deflection type screen

Graticule, waveform data,  menu,  setting condition,  title

Digit storage, two channels (A,B),  501 points of data for horizontal axis

Screen data can be hard-copied or printed via the GP-IB interface

Calibrates all level CAL1  CAL2  &FRQ CAL function

Calibration Total gain deflection   &log linearity

Calibration resolution bandwidth switching deviation, reference level deviation.

6 settings are saved to / recalled from memory

Residual response:

Marker:

Normal:

Delta:

Video bandwidth:

Level unit:

Quasi-peak detection:

6dB Bandwidth :

Charge-time constant:

Discharge-time constant:

Display - time constant:

RF input:

Impedance:

Max. Input level:

Attenuation:

Switching accuracy:

Sweep:

Time:

Trigger:

Range:

CRT display:

Display items:

Display method:

Direct plotting:

Calibration:

All cal:

LEVEL.CAL1:

LEVEL.CAL2:

Function memory:

Internal memory:

EXT. Control:

Function GP-IB:

Interface:

EXT output:

Tracking output:

2nd oscillator output:

X, Y and Z output:

X-axis output:

Y-axis output:

Z-axis output:

Check signal output:

Frequency:

Level Output:

Impedance:

IF output:

VIDEO output:

Dimensions & weight:

Power supply:

EMC EMI standard:

All function can be the externally controlled, except the power switch , CRT intensity, PMC control   GP-IB address, direct plotting.

SH1  AH1  T6  L4  SR1  RL1  PP0  DC1  DT1  C0

1st oscillator Output : 2.5214~4.7214GHz

2.5GHz

From 0V (left edge) to 10V (right edge)

From 0V (lower edge) to 1V (upper edge)

TTL level

50MHz  (50Ω)    -2dBm  (approx)

-2dBm(approx.)

50Ω

3.6MHz  ( 50Ω)    0dBm(at reference level line on CRT)

Composite and separate video output

451mm×320mm×177mm      20kg

50Hz±5%    220V±10%

GB6833.3 1-87  GB6833.4 9-87  GB6833 10-87

- digital technology and synthesized local oscillator supports high-speed/high-accuracy measurement.

- wide Freq. Range (9KHz~2200MHz) high resolution bandwidth (30Hz~1MH) low noise side band≤-100dBc(offset 20KHz), excellent dynamic range.

- EMI measurement function built-in quastpeak detection provide three bandwidth and three time constants based on CISPR

- one-touch hard copy or printout Full GP-IB control

- AUT-CAL itself through built-in signal source, each time the CAL key is pressed and the CAL data is processed at high-speed microprocessor and compensation. It is resulting level measurement accuracy of 1dB

The QF4032 is a portable spectrum analyzer with advanced functions,  covering 9KHz ~ 2200MHz.  It is most useful in accurately determining the signal frequency distribution,  spectrum analysis of waveform and signal level measurement.

It can be used in the analysis of the signal’s purity  (harmonics,  phase noise,  spurious……) and measurement of modulated signals (AM wave, FM wave, PM signals) and measurement of electromagnetic interference (EMI). It can also be used in the measurement of modern mobile communications (bandwidth occupied,  adjacent channel power,  parasitical signals……) and the detection of satellites communications,  CATV system.  Meanwhile, it can detect the quality of electronic devices such as VCO,  RF module,  Mixers  filter and ICs.

# Overview

# Features

# Main specifications

Page 4-1

Page 4-2

Page 4-3

Page 4-4